## Slide 1: Slicer’s Coordinate Systems
Chand T. John
Neuromuscular Biomechanics Lab
Stanford University

## Slide 2: What is a Coordinate System?
v1
v2
v3
O
A coordinate system for a three-dimensional space is a point O called the origin along with three linearly independent vectors v1, v2, and v3.
Linearly independent here means no two of the vectors are parallel and the three vectors do not all lie in the same plane.

## Slide 3: The RAS Coordinate System
Every scanned object has its own fixed
coordinate system known as the RAS
coordinate system. “RAS” stands for Right,
Anterior, Superior. For CT and MRI
machines, the RAS coordinate system is fixed
with respect to the scan table and the object or
subject being scanned (the “scanee” if you wish).

## Slide 4: The RAS Coordinate System
Right
Anterior
Superior
Inferior
Posterior
Left
R
A
S

## Slide 5: The Three Viewing Directions
R
A
S
L
P
I
A
L
R
P
S
L
R
I
S
P
A
I
Axial
Coronal
Sagittal
Axial
Coronal
Sagittal

## Slide 6: Medical Image Scan Directions
R
A
S
L
P
I
Axial: SI
Axial: IS
Sagittal: RL
Sagittal: LR
Coronal: AP
Coronal: PA
A medical scanner acquires data by scanning individual 2D slices along an axis:

## Slide 7: The IJK Coordinate System
Every medical scanner has its own coordinate
system called the IJK coordinate system. The
IJK coordinate system represents the actual
rectangular prism of data that is scanned,
instead of the position of the scan table (the
coordinate system of the “scanner”).

## Slide 8: The IJK Coordinate System
i
j
k
-k
-j
-i

## Slide 9: Coordinate System Conversion
We frequently need to convert between
coordinate systems, such as RAS and IJK.
The mapping from one coordinate system to
another is a 3D affine transformation: a
sequence of transformations consisting of a
shear, a reflection, a rotation, a scaling, and a
translation.

## Slide 10: Affine Transformations
x’
y’
z’
=
Translation
Reflection,
Rotation
Scaling
Shear
+
x
y
z
An affine transformation is simply a matrix multiplication applied to a point followed by a translation. Any affine transformation can be decomposed as follows:

## Slide 11: Affine Transformations
X
Z
Y
Let us examine how an affine transformation works within the standard coordinate system in 3D: this is the coordinate system with its origin at the point (0, 0, 0) and the three basis vectors (1, 0, 0), (0, 1, 0), and (0, 0, 1). We conventionally call label the axes representing these basis vectors as the X, Y, and Z axes.

## Slide 12: Shear
A shear transformation consists of shearing in each of the three coordinate planes: the XY, XZ, and YZ planes.
Thus z = 0 plane does not move at all
A shear in the XY plane translates every plane parallel to the XY plane by an amount in x and y proportional to the z value of that plane.

## Slide 13: Shear
x’
y’
z’
x
y
z
=
x’ = x + gxy
y’ = y
z’ = z + gzy
1 gx 0
0 1 0
0 gz 1
Shearing in the XZ plane:
Shearing in the YZ plane:
x’
y’
z’
x
y
z
=
x’ = x
y’ = y + hyx
z’ = z + hzx
1 0 0
hy 1 0
hz 0 1
x’
y’
z’
x
y
z
=
x’ = x + fxz
y’ = y + fyz
z’ = z
1 0 fx
0 1 fy
0 0 1
Shearing in the XY plane:

## Slide 14: Reflection
Reflection about the XY plane is a simple transformation: simply negate the Z coordinate.
Z
x’
y’
z’
x
y
z
=
1 0 0
0 1 0
0 0 -1

## Slide 15: Reflection
The reflection about the XY plane is the only reflection needed to generate any affine transformation, since a reflection about any other plane can be performed by conjugating the reflection about the XY plane with a translation and a rotation (and rotations and translations are already used in affine transformations):
Reflection about target plane
Reflection about XY plane
Translation, then rotation of target plane to XY plane
Rotation, then translation of XY plane to target plane
=

## Slide 16: Rotation
X
Y
Z
Every rotation in 3D is a combination of rotations around each of the standard coordinate axes.




## Slide 17: Rotation
x’
y’
z’
x
y
z
=
cos  0 sin 
0 1 0
-sin  0 cos 
Rotation around the Y axis:
Rotation around the Z axis:
x’
y’
z’
=
x’
y’
z’
x
y
z
=
1 0 0
0 cos  -sin 
0 sin  cos 
Rotation around the X axis:
x
y
z
cos  -sin  0
sin  cos  0
0 0 1

## Slide 18: Scaling
x’
y’
z’
x
y
z
=
sx 0 0
0 sy 0
0 0 sz
Scaling, also known as dilation, simply stretches the 3D space in each of the three standard coordinate directions X, Y, and Z by factors of Sx, Sy, and Sz respectively.
If Sx = Sy = Sz, then the scaling is said to be isotropic. Otherwise, the scaling is anisotropic.

## Slide 19: Translation
x’
y’
z’
x
y
z
=
tx
ty
tz
+
Translation is a very simple operation: it simply slides every point in space by a fixed vector (tx, ty, tz):

## Slide 20: Affine Transformations
We can see that all of the components of an affine transformation are either a multiplication by a 3 x 3 matrix or an addition of a 3 x 1 matrix. In fact, all transformations consisting of such a matrix multiplication followed by such a translation are affine transformations.
x’
y’
z’
x
y
z
=
tx
ty
tz
+
a11 a12 a13
a21 a22 a23
a31 a32 a33

## Slide 21: Homogeneous Coordinates
x’
y’
z’
1
x
y
z
1
=
a11 a12 a13 tx
a21 a22 a23 ty
a31 a32 a33 tz
0 0 0 1
When using homogeneous coordinates, we assign an extra coordinate of “1” to all points and an extra coordinate of “0” to all vectors:
3
2
6
3
2
6
1
4
7
3
4
7
3
0
Point
Vector
Using homogeneous coordinates, affine maps are written with a single matrix:

## Slide 22: Affine Transformation Summary
Any time you need to map one coordinate system to another, you are using an affine transformation.
Affine transformations are represented in Slicer using homogeneous coordinates, as 4 x 4 matrices.
Every affine transformation consists of some, if not all, of the components listed in the arrow below, applied in that order.
Shear, reflection, rotation,
scaling, translation
v1
v2
v3
O
v2’
O’
v1’
v3’

## Slide 23: The IJK Coordinate Systems
For each scan order (LR, RL, AP, PA, IS, SI),
there is a unique IJK coordinate system
associated with it. In Slicer, the resulting
transformation from the RAS coordinate
system to the IJK coordinate system is an
affine transformation consisting of a possible
reflection, a rotation by a multiple of 90o in zero
or more directions, a scaling, and a translation.

## Slide 24: The IJK Coordinate Systems
-i
j
k
i
-j
-k
k
-i
j
-k
i
-j
-i
k
j
i
-k
-j
-i
j
k
i
-j
-k
Axial: SI
Sagittal: RL
k
-i
j
-k
i
-j
Coronal: AP
-i
k
j
i
-k
-j
Axial: IS
Sagittal: LR
Coronal: PA

## Slide 25: RAS to IJK Transformation
So the RAS to IJK transformation typically contains a possible reflection, a rotation by a multiple of 90o in zero or more directions, a scaling, and a translation.
If a CT scanner’s scan table was tilted in one or more directions to limit radiation delivered to a patient’s vital organs, the RAS to IJK transformation will also contain an arbitrary rotation, not necessarily by a multiple of 90o. This scan table tilting is known as gantry tilt.
Sometimes the desired anatomy is imaged in an oblique direction to improve accuracy. This oblique acquisition will result in the RAS to IJK transformation containing a shear.
Thus, an arbitrary RAS-to-IJK transformation can contain all the components of an arbitrary affine transformation!
An RAS-to-IJK transformation should always be invertible: inverting the RAS-to-IJK 4 x 4 matrix yields the IJK-to-RAS matrix.

## Slide 26: RAS to IJK Transformation
Fortunately, the RAS-to-IJK transformation for every data set loaded into Slicer is
stored in Slicer. To obtain the 4 x 4 transformation matrix:
  Find the ID number of the data volume whose RAS-to-IJK transformation you want to obtain: go to the EditorVolumes and select the original grayscale to be volume you want. Then in the Tcl/Tk window, type puts [EditorGetOriginalID] and press Enter. The ID number of the volume you wanted is the output. We’ll pretend the ID number is 1.
  In the Tcl/Tk window, type the following commands: vtkMatrix4x4 mtx mtx DeepCopy [Volume(1,node) GetRasToIjk]
  Now mtx is a copy of the RAS-to-IJK matrix of the volume.
  To invert the matrix, type mtx Invert. mtx is now a copy of the IJK-to-RAS matrix for the volume.
  To apply mtx to the point 44 36 127 for example, simply type: mtx MultiplyPoint 44 36 127 1
  To print out the contents of mtx, type (without pressing Enter in between lines): for {set i 0} {$i < 4} {incr i} { for {set j 0} {$j < 4} {incr j} { puts [mtx GetElement $i $j] } }

## Slide 27: The XYS Coordinate System
In Slicer, the axial, sagittal, and coronal views each have their own XYS coordinate system.
In each window, each voxel on each image slice has its own (X, Y) screen coordinates and the slice number in which it lies.
Pixel (X, Y)
Slice S

## Slide 28: XYS to IJK Transformation
For the slice direction the user chose to segment in, we must first figure out how the XYS coordinates of any voxel map to the actual IJK coordinates of that voxel.
A clarification: the I coordinate labels the horizontal position of a voxel (ranges usually from 0 to 255, or 0 to 511). The J coordinate labels the vertical position of a voxel, and typically has the same range as the I coordinate. The K coordinate varies with the slice number in which a voxel lies.
Another clarification: for some reason, Slicer is designed so that in the axial and sagittal views, the slice number increases in the direction away from the viewer, while in the coronal view, the slice number increases in the direction toward the viewer. This is why the XYS axes are positioned differently for the coronal view when compared with the axial and sagittal views. The XYS axes also vary with scan order.
The XYS and IJK coordinate systems have their origins and whole numbered coordinates at the corners of voxels. The middle of a voxel will have non-integer coordinates that are multiples of ½.
In the following slides, the k=0 or k=n labels next to the IJK origin indicate whether the IJK origin is on slice 0, or on slice n, where n is the number of slices in the image data.

## Slide 29: Inaccuracies
There are some inaccuracies in my
explanations of XYS to IJK transformations.
The IJK to RAS transformations are likely to
be accurate. Thus the XYS to RAS
transformations are also inaccurate. However,
the Slicer code for transforming between
coordinate systems is accurate.

## Slide 30: Axial: IS IJK Coordinate System
-i
j
k
i
-j
-k
R
A
S
L
P
I
j
i
k
i
k
-j
Axial
Coronal
Sagittal
-k
j
i
x
y
s
x
y
s
x
y
s
i
k
j
i
k
j
k
i
j
i = -s
j = -x
k = y
i = x
j = s
k = y
i = x
j = y
k = s
k = 0

## Slide 31: Axial: SI IJK Coordinate System
-i
j
k
i
-j
-k
R
A
S
L
P
I
j
i
-k
i
-k
-j
Axial
Coronal
Sagittal
k
j
i
x
y
s
x
y
s
x
y
s
i
k
j
k
i
j
i = -s
j = -x
k = -y
i = x
j = s
k = -y
i = x
j = y
k = -s
i
k
j
k = n

## Slide 32: Sagittal: LR IJK Coordinate System
R
A
S
L
P
I
Axial
Coronal
Sagittal
k
-i
j
-k
i
-j
-i
-k
-j
j
-k
-i
j
i
-k
x
y
s
x
y
s
x
y
s
i
j
k
i
j
k
j
i
k
i = x
j = y
k = s
i = -s
j = y
k = -x
i = -y
j = s
k = -x
k = 0

## Slide 33: Sagittal: RL IJK Coordinate System
R
A
S
L
P
I
Axial
Coronal
Sagittal
k
-i
j
-k
i
-j
-i
k
-j
j
k
-i
j
i
k
x
y
s
x
y
s
x
y
s
i
j
k
j
i
k
i = x
j = y
k = -s
i = -s
j = y
k = x
i = -y
j = s
k = x
i
j
k
k = n

## Slide 34: Coronal: PA IJK Coordinate System
R
A
S
L
P
I
Axial
Coronal
Sagittal
-i
k
j
i
-k
-j
k
i
-j
j
i
k
j
-k
i
x
y
s
x
y
s
x
y
s
i
j
k
i
j
k
j
i
k
i = -s
j = y
k = -x
i = x
j = y
k = s
i = x
j = s
k = y
k = 0

## Slide 35: Coronal: AP IJK Coordinate System
R
A
S
L
P
I
Axial
Coronal
Sagittal
-i
k
j
i
-k
-j
-k
i
-j
j
i
-k
j
k
i
x
y
s
x
y
s
x
y
s
i
j
k
j
i
k
i = -s
j = y
k = x
i = x
j = y
k = -s
i = x
j = s
k = -y
i
j
k
k = n

## Slide 36: XYS to IJK Transformations
So, for Axial: IS images:
i
j
k
x
y
s
=
=
1 0 0
0 1 0
0 0 1
x
y
s
Axial XYS to IJK matrix is:
Sagittal XYS to IJK matrix is:
i
j
k
-s
-x
y
=
=
0 0 -1
-1 0 0
0 1 0
x
y
s
Coronal XYS to IJK matrix is:
i
j
k
x
s
y
=
=
x
y
s
1 0 0
0 0 1
0 1 0

## Slide 37: XYS to IJK Transformations
For Axial: SI images:
i
j
k
x
y
-s
=
=
1 0 0
0 1 0
0 0 -1
x
y
s
Axial XYS to IJK matrix is:
Sagittal XYS to IJK matrix is:
i
j
k
-s
-x
-y
=
=
0 0 -1
-1 0 0
0 -1 0
x
y
s
Coronal XYS to IJK matrix is:
i
j
k
x
s
-y
=
=
x
y
s
1 0 0
0 0 1
0 -1 0

## Slide 38: XYS to IJK Transformations
For Sagittal: LR images:
Axial XYS to IJK matrix is:
Sagittal XYS to IJK matrix is:
i
j
k
-y
s
-x
=
=
0 -1 0
0 0 1
-1 0 0
x
y
s
Coronal XYS to IJK matrix is:
i
j
k
-s
y
-x
=
=
x
y
s
0 0 -1
0 1 0 -1 0 0
i
j
k
x
y
s
=
=
1 0 0
0 1 0
0 0 1
x
y
s

## Slide 39: XYS to IJK Transformations
For Sagittal: RL images:
Axial XYS to IJK matrix is:
Sagittal XYS to IJK matrix is:
i
j
k
-y
s
x
=
=
0 -1 0
0 0 1
1 0 0
x
y
s
Coronal XYS to IJK matrix is:
i
j
k
-s
y
x
=
=
x
y
s
0 0 -1
0 1 0
1 0 0
i
j
k
x
y
-s
=
=
1 0 0
0 1 0
0 0 -1
x
y
s

## Slide 40: XYS to IJK Transformations
For Coronal: PA images
Axial XYS to IJK matrix is:
Sagittal XYS to IJK matrix is:
i
j
k
-s
y
-x
=
=
0 0 -1
0 1 0
-1 0 0
x
y
s
Coronal XYS to IJK matrix is:
i
j
k
x
s
y
=
=
x
y
s
1 0 0
0 0 1
0 1 0
i
j
k
x
y
s
=
=
1 0 0
0 1 0
0 0 1
x
y
s

## Slide 41: XYS to IJK Transformations
For Coronal: AP images
Axial XYS to IJK matrix is:
Sagittal XYS to IJK matrix is:
i
j
k
-s
y
x
=
=
0 0 -1
0 1 0
1 0 0
x
y
s
Coronal XYS to IJK matrix is:
i
j
k
x
s
-y
=
=
x
y
s
1 0 0
0 0 1
0 -1 0
i
j
k
x
y
-s
=
=
1 0 0
0 1 0
0 0 -1
x
y
s

## Slide 42: Medical Image Parameters
Scan order
Pixel spacing in horizontal direction (mm)
Pixel spacing in vertical direction (mm)
Slice thickness (mm)
Origin coordinates in RAS space (mm, mm, mm)

## Slide 43: Medical Image Parameters
Pixel spacing in horizontal direction (x)
Pixel spacing in vertical direction (y)
Slice thickness (z)
Origin coordinates in RAS space (Ox, Oy, Oz)
z
x
y
(Ox, Oy, Oz)
2D Images

## Slide 44: Transformation Example
We will now demonstrate a detailed example of the possible XYS to IJK and IJK to RAS transformations that can occur.
This particular example has no gantry tilt or oblique acquisition: so the transformations here contain only reflection, rotation by multiples of 90o, scaling, and translation.

## Slide 45: IJK to RAS Transformations
For Axial: IS images:
Note: the RasToIjk matrix in Slicer that we are using is that of the image, and not of the labelmap that represents the segmentation. We invert RasToIjk to get the IjkToRas matrix.
R
A
S
-x * i + Ox
y * j + Oy
z * k + Oz
=
=
i
j
k
-x 0 0 Ox
0 y 0 Oy
0 0 z Oz
0 0 0 1
-i
j
k
i
-j
-k
R
A
S
L
P
I
Axial: IS IJK
RAS
(Ox, Oy, Oz)
+
Ox
Oy
Oz
Or in Slicer’s homogeneous coordinates:
R
A
S
1
=
-x 0 0
0 y 0
0 0 z
i
j
k
1

## Slide 46: IJK to RAS Transformations
For Axial: SI images:
R
A
S
-x * i + Ox
y * j + Oy
-z * k + Oz
=
=
i
j
k
-x 0 0 Ox
0 y 0 Oy
0 0 -z Oz
0 0 0 1
-i
j
k
i
-j
-k
R
A
S
L
P
I
Axial: SI IJK
RAS
(Ox, Oy, Oz)
+
Ox
Oy
Oz
Or in Slicer’s homogeneous coordinates:
R
A
S
1
=
-x 0 0
0 y 0
0 0 -z
i
j
k
1

## Slide 47: IJK to RAS Transformations
For Sagittal: LR images:
R
A
S
z * k + Ox
-x * i + Oy
y * j + Oz
=
=
i
j
k
0 0 z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
R
A
S
L
P
I
Sagittal: LR IJK
RAS
(Ox, Oy, Oz)
+
Ox
Oy
Oz
Or in Slicer’s homogeneous coordinates:
R
A
S
1
=
0 0 z
-x 0 0
0 y 0
i
j
k
1
k
-i
j
-k
i
-j

## Slide 48: IJK to RAS Transformations
For Sagittal: RL images:
R
A
S
-z * k + Ox
-x * i + Oy
y * j + Oz
=
=
i
j
k
0 0 -z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
R
A
S
L
P
I
Sagittal: RL IJK
RAS
(Ox, Oy, Oz)
+
Ox
Oy
Oz
Or in Slicer’s homogeneous coordinates:
R
A
S
1
=
0 0 -z
-x 0 0
0 y 0
i
j
k
1
k
-i
j
-k
i
-j

## Slide 49: IJK to RAS Transformations
For Coronal: PA images:
R
A
S
-x * i + Ox
y * j + Oz
z * k + Oy
=
=
i
j
k
-x 0 0 Ox
0 0 z Oy
0 y 0 Oz
0 0 0 1
R
A
S
L
P
I
Coronal: PA IJK
RAS
(Ox, Oy, Oz)
+
Ox
Oy
Oz
Or in Slicer’s homogeneous coordinates:
R
A
S
1
=
-x 0 0
0 0 z
0 y 0
i
j
k
1
-i
k
j
i
-k
-j

## Slide 50: IJK to RAS Transformations
For Coronal: AP images:
R
A
S
-x * i + Ox
y * j + Oz
-z * k + Oy
=
=
i
j
k
-x 0 0 Ox
0 0 -z Oy
0 y 0 Oz
0 0 0 1
R
A
S
L
P
I
Coronal: AP IJK
RAS
(Ox, Oy, Oz)
+
Ox
Oy
Oz
Or in Slicer’s homogeneous coordinates:
R
A
S
1
=
-x 0 0
0 0 -z
0 y 0
i
j
k
1
-i
k
j
i
-k
-j

## Slide 51: XYS to IJK to RAS
Axial: IS
-x 0 0 Ox
0 y 0 Oy
0 0 z Oz
0 0 0 1
1 0 0 0
0 1 0 0
0 0 1 0
0 0 0 1
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
0 0 -1 0
-1 0 0 0
0 1 0 0
0 0 0 1
Coronal XYS to RAS matrix is:
1 0 0 0
0 0 1 0
0 1 0 0
0 0 0 1
=
-x 0 0 Ox
0 y 0 Oy
0 0 z Oz
0 0 0 1
-x 0 0 Ox
0 y 0 Oy
0 0 z Oz
0 0 0 1
-x 0 0 Ox
0 y 0 Oy
0 0 z Oz
0 0 0 1
=
=
0 0 x Ox
-y 0 0 Oy
0 z 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 y Oy
0 z 0 Oz
0 0 0 1
XYS to IJK
IJK to RAS
XYS to RAS

## Slide 52: XYS to IJK to RAS
Axial: SI
-x 0 0 Ox
0 y 0 Oy
0 0 -z Oz
0 0 0 1
1 0 0 0
0 1 0 0
0 0 -1 0
0 0 0 1
0 0 -1 0
-1 0 0 0
0 -1 0 0
0 0 0 1
1 0 0 0
0 0 1 0
0 -1 0 0
0 0 0 1
=
-x 0 0 Ox
0 y 0 Oy
0 0 z Oz
0 0 0 1
-x 0 0 Ox
0 y 0 Oy
0 0 -z Oz
0 0 0 1
-x 0 0 Ox
0 y 0 Oy
0 0 -z Oz
0 0 0 1
=
=
0 0 x Ox
-y 0 0 Oy
0 z 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 y Oy
0 z 0 Oz
0 0 0 1
XYS to IJK
IJK to RAS
XYS to RAS
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:

## Slide 53: XYS to IJK to RAS
Sagittal: LR
0 -1 0 0
0 0 1 0
-1 0 0 0
0 0 0 1
1 0 0 0
0 1 0 0
0 0 1 0
0 0 0 1
0 0 -1 0
0 1 0 0
-1 0 0 0
0 0 0 1
=
-z 0 0 Ox
0 x 0 Oy
0 0 y Oz
0 0 0 1
=
=
0 0 z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
-z 0 0 Ox
0 0 x Oy
0 y 0 Oz
0 0 0 1
XYS to IJK
IJK to RAS
XYS to RAS
0 0 z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
0 0 z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
0 0 z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:

## Slide 54: XYS to IJK to RAS
Sagittal: RL
0 -1 0 0
0 0 1 0
1 0 0 0
0 0 0 1
1 0 0 0
0 1 0 0
0 0 -1 0
0 0 0 1
0 0 -1 0
0 1 0 0
1 0 0 0
0 0 0 1
=
-z 0 0 Ox
0 x 0 Oy
0 0 y Oz
0 0 0 1
=
=
0 0 z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
-z 0 0 Ox
0 0 x Oy
0 y 0 Oz
0 0 0 1
XYS to IJK
IJK to RAS
XYS to RAS
0 0 -z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
0 0 -z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
0 0 -z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:

## Slide 55: XYS to IJK to RAS
Coronal: PA
1 0 0 0
0 0 1 0
0 1 0 0
0 0 0 1
0 0 -1 0
0 1 0 0
-1 0 0 0
0 0 0 1
1 0 0 0
0 1 0 0
0 0 1 0
0 0 0 1
=
-x 0 0 Ox
0 z 0 Oy
0 0 y Oz
0 0 0 1
=
=
0 0 x Ox
-z 0 0 Oy
0 y 0 Oz
0 0 0 1
XYS to IJK
IJK to RAS
XYS to RAS
-x 0 0 Ox
0 0 z Oy
0 y 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 z Oy
0 y 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 z Oy
0 y 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 z Oy
0 y 0 Oz
0 0 0 1
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:

## Slide 56: XYS to IJK to RAS
Coronal: AP
1 0 0 0
0 0 1 0
0 -1 0 0
0 0 0 1
0 0 -1 0
0 1 0 0
1 0 0 0
0 0 0 1
1 0 0 0
0 1 0 0
0 0 -1 0
0 0 0 1
=
-x 0 0 Ox
0 z 0 Oy
0 0 y Oz
0 0 0 1
=
=
0 0 x Ox
-z 0 0 Oy
0 y 0 Oz
0 0 0 1
XYS to IJK
IJK to RAS
XYS to RAS
-x 0 0 Ox
0 0 -z Oy
0 y 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 -z Oy
0 y 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 -z Oy
0 y 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 z Oy
0 y 0 Oz
0 0 0 1
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:

## Slide 57: Summary: XYS to RAS
Axial: IS and Axial: SI
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:
-x 0 0 Ox
0 y 0 Oy
0 0 z Oz
0 0 0 1
0 0 x Ox
-y 0 0 Oy
0 z 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 y Oy
0 z 0 Oz
0 0 0 1

## Slide 58: Summary: XYS to RAS
Sagittal: LR and Sagittal: RL
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:
-z 0 0 Ox
0 x 0 Oy
0 0 y Oz
0 0 0 1
0 0 z Ox
-x 0 0 Oy
0 y 0 Oz
0 0 0 1
-z 0 0 Ox
0 0 x Oy
0 y 0 Oz
0 0 0 1

## Slide 59: Summary: XYS to RAS
Coronal: PA and Coronal: AP
Axial XYS to RAS matrix is:
Sagittal XYS to RAS matrix is:
Coronal XYS to RAS matrix is:
-x 0 0 Ox
0 z 0 Oy
0 0 y Oz
0 0 0 1
0 0 x Ox
-z 0 0 Oy
0 y 0 Oz
0 0 0 1
-x 0 0 Ox
0 0 z Oy
0 y 0 Oz
0 0 0 1

## Slide 60: XYS to RAS in Slicer
Because of the complexity of transforming XYS points to RAS, Slicer provides convenience routines for converting from XYS to RAS in real-time: meaning, it will only convert points that are currently lying on a visible slice in either the axial, sagittal, or coronal view windows.
The following C++ code segment converts a point screenPt containing just X and Y screen coordinates to an RAS point rasPt. The slice window containing screenPt is determined using the GetActiveSlice method. This code is valid only in the files vtkMrmlSlicer.h and vtkMrmlSlicer.cxx: int screenPt[2]; screenPt[0] = 36; screenPt[1] = 44; int as = this->GetActiveSlice(); vtkFloatingPointType rasPt[3]; this->SetReformatPoint(as, screenPt[0], screenPt[1]); this->GetWldPoint(rasPt); // Now rasPt contains the RAS coordinates of the XYS // point screenPt

## Slide 61: XYS to RAS in Slicer
In any Tcl file, the following code will convert a screen point (x, y) in the active 2D window to an RAS point (xwld, ywld, zwld): set s $Interactor(s) # gets active slice window Slicer SetReformatPoint $s $x $y scan [Slicer GetWldPoint] “%g %g %g” xwld ywld zwld
For converting from (x, y) to an IJK point (xijk, yijk, zijk) in Tcl: set s $Interactor(s) # gets active slice window Slicer SetReformatPoint $s $x $y scan [Slicer GetIjkPoint] “%g %g %g” xijk yijk zijk set xijk [expr int($xijk)] set yijk [expr int($yijk)] set zijk [expr int($zijk)]

## Slide 62: Position Matrix
Just like GetRasToIjk retrieves a 4 x 4 RAS-to-IJK matrix for a volume, GetPosition retrieves a 4 x 4 matrix representing the position of the RAS coordinate system relative to the IJK coordinate system.
My conjecture is that the position matrix contains only the rigid transformation (reflection, rotation, translation) of the RasToIjk matrix. More precisely, in an experiment I found that: IjkToRas = (Translate) * (Rotate/reflect) * (Scale) Position = (Translate) * (Rotate/reflect)